The BRAHMS Experiment at RHIC

I. G. Bearden7, D. Beavis1, C. Besliu10, Y. Blyakhman6, J. Brzychczyk4, 
B. Budick6, H. Bøggild7, C. Chasman1, C. H. Christensen7, P. Christiansen7, 
J. Cibor3, R. Debbe1, J. J. Gaardhøje7, K. Grotowski4, K. Hagel8, O. Hansen7, 
A. Holm7, A. K. Holme12, H. Ito11, E. Jakobsen7, A. Jipa10, J. I. Jørdre9, 
F. Jundt2, C. E. Jørgensen7, T. Keutgen8, E. J. Kim1,5, T. Kozik4, T. M. Larsen12, J. H. Lee1, Y. K. Lee5, G. Løvhøiden12, Z. Majka4, A. Makeev8, B. McBreen1, M. Murray8, J. Natowitz8, B. S. Nielsen7, K. Olchanski1, J. Olness1, 
D. Ouerdane7, R. Płaneta4, F. Rami2, D. Röhrich9, B. H. Samset12, 
S. J. Sanders11, R. A. Sheetz1, Z. Sosin4, P. Staszel7, T. F. Thorsteinsen9+, 
T. S. Tveter12, F. Videbæk1, R. Wada8, A. Wieloch4, and  I. S. Zgura10

(BRAHMS Collaboration)

- Brookhaven National Laboratory, New York, 11973, USA.

- Institut de Recherches Subatomiques and Université Louis Pasteur, Strasbourg, France.

- Institute of Nuclear Physics, Krakow, Poland

- Jagiellonian University, Krakow, Poland.

- Johns Hopkins University, Baltimore, Maryland 21218, USA.

- New York University, New York, New York 10003, USA.

- Niels Bohr Institute, University of Copenhagen,.Denmark

- Cyclotron Institute, Texas A&M University, College Station, Texas 77843, USA

- University of Bergen, Department of Physics, Bergen, Norway.

- University of Bucharest, Romania.

- University of Kansas, Lawerence, Kansas 66045, USA.

- University of Oslo, Department of Physics, Oslo, Norway.

+ Deceased

version 26-Sep-2001-1

Abstract. The BRAHMS experiment at RHIC was conceived to pursue the understanding of nuclear matter under extreme conditions by detailed measurements of charged hadrons over the widest possible range of rapidity and transverse momentum.  The experiment consists of two spectrometers with complementary charged hadron detection capabilities as well as a series of global detectors for event characterization.  A series of tracking detectors, time of flight arms and Čerenkov detectors enables momentum determination and particle identification over a wide range of rapidity and transverse momentum.  Technical details and performance results are presented for the various detector subsystems.  The performance of the system as a whole is shown to meet the goals of the experiment.

Introduction

The BRAHMS experiment was first proposed in 1991 as a small experiment at RHIC, and subsequently approved for construction that started in 1997.  The experiment was partly operational for the first RHIC beams in July 2000 and is essentially complete for the run in 2001.  The BRAHMS experiment pursues the understanding of properties of nuclear matter by detailed and precise measurements of charged hadrons over a large range of rapidity and transverse momentum.  The present paper describes the design considerations for the experiment, the layout, construction and details of the individual detector components and finally presents some performance results.

# BRAHMS Physics GOALS

The BRAHMS detectors are designed to cover the widest possible range in polar angle relative to the beam direction with precise momentum determination and good particle identification (PID).  In this way , the BRAHMS experiment will be complementary to the other three RHIC experiments that emphasize reaction angles near 90 degrees, the mid rapidity region.  The momentum span from 90 degrees to very forward angles (here 2.3 degrees) is from a few hundred MeV/c to 25-30 GeV/c, a condition that led to the use of two separate spectrometers, one for low to medium momenta and one for high momenta, both with comparatively small solid angles.  The actual sizes of the spectrometers are largely determined by practical considerations, such as the size and geometry of the experimental hall.  The maximum forward angle was set by the size and location of the RHIC DX magnets that control the beam intersection region.  The BRAHMS spectrometer concept is an extension of the spectrometers used in AGS experiments E802 and E866 [1,2].

In characterizing particle spectroscopy at relativistic energies it is customary to use kinematic variables that have simple properties under Lorentz transformation along the beam direction (boost transformations), rather than polar angle and momentum.  Here we shall use two variables to describe the phase space of the observed particles: transverse momentum (pt) and rapidity (y).  pt is defined as pt=psinθ where p is the magnitude of the momentum and θ the polar angle with respect to the beam direction; pt is invariant under boost transformations.  Rapidity is defined as y=12ln((E+p||)(E−p||)), where E is the energy and p||=pcosθ.  Rapidity is additive under boost transformations and approaches velocity divided by the vacuum velocity of light in the non-relativistic limit.

Figure 1 shows the y-pt acceptance for p, K and protons in BRAHMS for the different spectrometer settings.  Region I and II in the figure show the regions accessible by the forward (high momentum) spectrometer (FS) and region III the range accessible to the mid rapidity (low momentum) spectrometer (MRS).  It should be noted that measurements in region III can be made simultaneously with those in any of the first two regions.

The rapidity coverage of the entire spectrometer is from y0 to y4.  For comparison the beam rapidity is 5.37 at the maximum RHIC collision energy of 100 GeV/u Au + Au.  The other RHIC experiments are limited to |y|<2 for particle spectroscopy with good PID.  The PID in BRAHMS utilizes standard spectrometer techniques, time of flight (two systems in the FS and one in MSR), threshold Čerenkov and ring imaging Čerenkov detectors (both in FS).  The pt coverage is typical for soft collisions, but approaches the hard collision region  (i.e. y ~ 2-3 and pt up to 3-4 GeV/c) in region I.

Both spectrometers have small solid angles, 0.8 msr for the FS and 5.6 msr for the MRS, so the broad rapidity coverage is obtained by systematically combining many polar angle and magnetic field settings of both spectrometers.  The BRAHMS detector measures averages over many events, and the events are characterized by a set of global detectors that can measure the charged particle multiplicity in a large solid angle near zero rapidity (the Tiles and Si detectors), the forward and backward charged multiplicities in a number of small solid angle Čerenkov detectors (the beam-beam counters) and finally neutrons near zero degrees with the zero degree calorimeters (ZDC).  The ZDC calorimeters are common to all RHIC experiments.

The detector design enables the BRAHMS experiment to do hadron spectroscopy in a large rapidity and pt range.  The first goal of the experiment is to map out the rapidity density distributions for protons and antiprotons, a sensitive measurement for the dynamics of heavy ion collisions at these high energies.  Other goals are the spectroscopy of pions and kaons as well as light clusters (e.g. deuterons and anti deuterons), all of which provide essential information on the thermodynamical characteristics of the collision processes.

Figure 1.  BRAHMS Acceptance for pions, kaons and protons.  Region I and II show the acceptance of the forward spectrometer with two different configurations, see below.  Region III shows the acceptance for the mid rapidity spectrometer.  The dotted curves marked 4, 5, 10, 25 and 35 GeV/c are curves of constant momentum, while the set marked 95, 30, 15, 5, 2 degrees indicate curves at constant polar angle. The fully drawn curves mark the acceptance borders.

# Experimental Layout

A perspective view of the BRAHMS detector layout is presented in Figure 2, where the two magnetic spectrometers, and the complementary global detectors are shown.  The Forward- Spectrometer (FS) is positioned to the right of and near the beampipe at 2.3 degrees and the Mid- Rapidity Spectrometer (MRS) to the left of the beam-pipe near 90 degrees.

## Forward Spectrometer

The forward spectrometer contains four dipole magnets, D1-D4, for sweeping and analyzing particles emerging from the reaction.  D1 is of a septum design in order to fit into the space near the beam-pipe at the most forward setting of ~2.3o.  The maximum field in D1 is B=1.2 T, enough to bend 25 GeV/c particles into the aperture of the subsequent magnet, D2.  The magnets, D2-D4, are of conventional design.  The forward spectrometer has a solid angle of 0.8 msr.

The first two tracking detectors, T1 and T2, are Time Projection Chambers (TPCs) which provide good three-dimensional track recognition and a high degree of background rejection in a high multiplicity environment.  The TPCs operate outside of the magnetic fields and have drift distances of ~20cm.  The remaining tracking detectors, T3-T5, are Drift Chambers (DCs) each consisting of three modules, where each module contains 10 planes for T3 and 8 planes for T4 and T5.  A photograph of the Forward Spectrometer is shown in Figure 3.

The spectrometer can be operated in two configurations: one for the highest momenta and one for intermediate momenta.  In the high momentum mode particles are deflected by the dipoles D1 and D2 into the back end of the spectrometer where the tracking is performed by T2 and the drift chambers, and the momenta are determined by combining the tracking and the fields of the dipole magnets D3 and D4.

In the intermediate momentum configuration the particles are tracked by T1 and T2 and the momenta are determined using the field setting of D2.  The D1-D2, T1-T2 part of the spectrometer, which constitutes the intermediate momentum configuration, is mounted on a platform that can move from 2.3 degrees to 30 degrees, while the high momentum configuration platform can move from 2.3 to 15 degrees.

Figure 2. The BRAHMS detectors in perspective.  The spectrometer magnets are denoted D1 through D5, the time of flight hodoscopes in FS are H1 and H2, and the tracking counters T1-T5, where T1 and T2 are time projection chambers and T3, T4 and T5 are multi-wire drift chambers.  The two Čerenkov counters in FS are C1 and RICH.  The tracking chambers of the MRS are TPM1 and TPM2, while the time of flight system is the TOFW.  The global detectors are denoted as Beam-Beam counters and Multiplicity.  The ZDCs (not depicted) are situated beyond the DX beam magnets.  The distance between the two DX beam magnets is 18m.

Particle identification in the forward spectrometer is based on two time of flight (TOF) hodoscopes, H1 and H2 and two Čerenkov detectors (C1 and RICH).  H1 and H2 consist of 40 and 32 plastic scintillator slats, respectively.  Each scintillator is instrumented with two phototubes, one on either end.  H1 is positioned 8.6m from the nominal vertex and H2 is positioned 19m from the nominal vertex.  The TOF hodoscopes are designed so that p/K and K/p separation can be achieved to momenta of 3.3 and 5.7 GeV/c, respectively, for H1 and to 5.8 and 8.5 GeV/c, respectively, for H2.  H1 and C1 are mounted on the intermediate momentum configuration platform, while the RICH detector is at the far end of the high momentum configuration.  The threshold Čerenkov detector, C1 extends the /K/p identification behind H1 to 9 GeV/c while the ring imaging Čerenkov detector (RICH) detector allows p/K/p separation in the high momentum configuration up to ~25 GeV/c.

Mid-Rapidity Spectrometer

Figure 3. The Forward Spectrometer photographed from the far end towards the interaction region.  The blue yokes of D3 and D4 appear in the middle of the picture.  The rails for the high momentum configuration’s movable platform can be seen.  The RICH detector is to the left in the picture, while at the right side the yellow yoke of D2 appears.

The Mid-Rapidity Spectrometer is a single dipole magnetic spectrometer with a geometric acceptance of 6.5 msr and it covers the angular range of 30o<θ<95o.  The particle identification requirements are simpler in the MSR than in the FS because the typical particle momentum at the larger angles is lower. Only time of flight is therefore used for this purpose.  The D5 magnet in the MRS is placed between the two TPCs and has an aperture of ~10x35 cm2 at a distance of 1.9 m from the nominal vertex.

Tracking in the MRS is done with the two TPCs (TPM1 and TPM2) which are similar in electrostatic design to the FS TPCs but different in general geometry.  Particle identification is accomplished with a highly segmented scintillator time of flight wall (TOFW) which provides p/K separation to momenta of 2 GeV/c and K/p separation to momenta of 3.5 GeV/c.

## Global Detectors

BRAHMS employs a series of detectors to enable event characterization by observation of certain overall features of each detected collision e.g. features that provide information on the collision centrality.  Some of the detectors also provide signals for the first level triggers

A two component multiplicity array surrounds the interaction vertex (marked “Multiplicity in Figure 2).  The inner array is made of Si strip detectors arranged as a hexagonal barrel around the beam axis.  The outer array of plastic scintillator tiles surrounds the silicon strip detectors, also mounted as a hexagonal barrel coaxial with the inner detector.  This hybrid system measures the energy loss of charged particles that traverse the detector surfaces.  Both arrays are segmented so angular distributions can be obtained.  By proper calibration of the Si and plastic detectors with minimum ionizing particles and use of detector simulation codes (GEANT [3]) it is possible to convert the energy loss data to event multiplicity within some systematic uncertainties, typically of the order 5-10%.  The quantity of interest is the multiplicity as function of pseudo rapidity, dN/d, where the pseudo rapidity is defined as η=12ln((p+ppar)(p−ppar))=−ln(tanθ2).Pseudo rapidity is a useful variable as it requires only knowledge of the polar angle and approaches rapidity at high momenta where momentum p approaches energy E.  Given the geometry of the detectors as shown in the photograph in Figure 4, the pseudo rapidity coverage is ||<2, corresponding to 15.4o<θ<164.6o, measured from the position of the nominal interaction vertex.

The beam-beam counters are located on either side of the interaction vertex 220 cm from the nominal interaction vertex and very close to the beam pipe (see Figure 2).  They are constructed from two different sized Čerenkov radiators glued to matching photo multiplier tubes.  The beam-beam counters are positioned at pseudo-rapidities of 3.1, 3.4 and 3.6 relative to the nominal interaction vertex point.  These counters provide a start time and a level 0 trigger.  The 50 ps time resolution allows determination of the actual interaction vertex position to an accuracy of better than 2cm from measurements of the difference in arrival times of the first particles to reach the two arrays.

The Zero Degree Calorimeters [4] (ZDCs) are placed behind the DX magnets on both sides of the interaction vertex 18m from the nominal interaction point.  They measure the total energy of neutrons in a cone of θ<2 mr around the downstream beam directions.  The ZDC resolution is 21% for 1 neutron at 65 GeV.  Coincidence events in the two ZDC detectors can provide a near minimum bias trigger, making it useful as an event trigger as well as a luminosity monitor [5].  In addition, the neutron multiplicity is also known to be correlated with the event geometry [6] and can therefore be used together with the data from the multiplicity array and the beam-beam counters to estimate the centrality of the collision.

Figure 4. Global detectors.  The photo shows the hexagonal barrel arrangement of the plastic tiles (dark square surfaces) , the beam line and to the right, just in front of the yellow yoke of the D1 magnet, one of the beam-beam counter arrays, where the radiators show as dark circles. The Si strip detectors have been removed for the picture and the other beam-beam counter array is outside the photo to the left

There are identical sets of ZDCs in all of the RHIC experiments.  They provide luminosity measurements for the RHIC machine group as well as neutron measurements.  Since the ZDCs provide some centrality information, the placement of identical detectors in all four of the RHIC experiments provides the possibility of comparing results from the different experiments obtained with nearly identical event characteristics.

# BRAHMS Subsystems

The various subsystems that make up BRAHMS are mounted on three movable platforms.  The MRS is mounted on a platform that turns around a pivot point located vertically under the nominal interaction point. The D1 and D2 magnets with two TPC detectors (T1 and T2) plus the H1 hodoscope and the C1 Čerenkov counter are mounted on a second pivoted platform, while the high momentum part of the FS sits on a third platform that moves on two concentric rails.  All angle settings can be remotely controlled to better than 0.08 mrad.  The internal geometry of the spectrometer components on a platform is constant to better than 0.1 mm under any allowed rotation.

In the following we discuss in detail the different subsystems which make up BRAHMS.  Magnets are discussed first, followed by the detector subsystems, triggering data acquisition and online monitoring.

## Magnets

BRAHMS has 5 dipole magnets, D1-D5, designed to provide momentum dispersion and determination for the charged particles produced in the heavy ion collisions.  The main requirements stem from the desire to cover as large an angular (rapidity) and momentum interval as possible with good momentum resolution.  Limitations to meeting these requirements come from the multiplicity capability of the tracking detectors as well as the overall costs of the experiment.  These requirements led to the design of a mid-rapidity spectrometer magnet, D5, to be used alone for the lower momentum particles expected at mid-rapidity and a forward spectrometer magnet array D1-D4, for providing the bend necessary for the higher momentum particles expected in the forward direction.  With the exception of D4, the magnets are of the window frame design with the return yoke primarily on one side.  This design was selected in order to secure the greatest forward angle coverage for the spectrometers.  Typically, the maximum operating fields are close to 1.2 T, limited mostly by the stray field allowed on the RHIC ring.

The dimensions of the return yoke have been established by requiring that the iron thickness be one half of the horizontal gap plus a coil width multiplied by the ratio of the maximum field to 1.2 T.  All magnets are thus slightly saturated at their maximum field.  The dimensions of the coils and the size of the conductor are fixed by the available space, required Ampere-turns, and the requirements for cooling.  The top and bottom coils are powered in series.

D1 is the most forward magnet and it is closest to the beam line.  These conditions imposed two important design criteria.  The first criterion was that the magnet be designed to fit into the space near the beam pipe when placed at the most forward angle of 2.3 degrees while delivering a field up to 1.3 T to give sufficient bending power to particles with momenta up to 25 GeV/c.  The second criterion was that the design result in fringe fields of D1 small enough not to affect the beam optics when D1 is placed at the most forward angle of 2.3 degrees with high field.

D1 has two “long” symmetric saddle (bed stand shaped) coils of 12 (3x4) turns each.  The cross section of the coils is 4x4cm2 and the dimensions of the magnet are given in Table 1.  The magnet is designed to bend a particle with unit charge and momentum of 23.6 GeV/c 1.95 degrees at 1.3 T. In order to minimize the distance between the magnet volume and the beam pipe (open side of the iron), the cross section of the iron is designed to be “C”-shaped and the side nearest the beam pipe (open side of the iron) is closed with a 0.95 cm thick “septum plate” made of the same material as the magnet iron to reduce the fringe field at the beam pipe. The design of the magnet was optimized by the use of detailed three dimensional field calculations using the code TOSCA [7].

Table 1.  BRAHMS magnet parameters.  All lengths are in cm.

|  | D1 | D2 | D3 | D4 | D5 |
| --- | --- | --- | --- | --- | --- |
| Gap dimensions |  |  |  |  |  |
| Height | 8.0 | 15.2 | 25.4 | 32.0 | 10.16 |
| Width | 20.0 | 27.3 | 40.0 | 50.8 | 35.6 |
| Length | 200 | 160 | 200 | 187 | 76.2 |
| Max. field [T] | 1.3 | 1.5 | 1.2 | 1.2 | 1.2 |
| B·dl [Tm] | 2.6 | 2.8 | 2.58 | 2.34 | 1.21 |
| Current [A] | 3200 | 3570 | 2813 | 2400 | 2210 |
| Magnet dimensions |  |  |  |  |  |
| Length | 225 | 200. | 260 | 244 | 99 |
| Width | 49. | 98. | 142. | 239. | 101.6 |
| Shield thick | 2.54 | 5.08 | 10.16 | 15.24 | 2.54 |

The field vector outside the gap of D1 on the beam side of the magnet was scanned at a distance of 10.7 cm away from the septum plate along the length of the gap.  The mean field in the gap in the vertical direction (y-direction) was 1.2 T, and the bending power of the y-component of the fringe field was measured in the scan to be B·dl=0.015 Tm.  The small strength of the fringe field has a negligible effect on the 100 GeV/c per nucleon beam.

TOSCA calculations were also used to calculate the fringe fields of D1.  The results of these calculations are shown in Figure 5 where we plot By versus x, the distance from the center of the magnet in the horizontal direction.  At 23 cm, the distance of the beam pipe from the center of D1 when the spectrometer is placed at 2.3 degrees, the fringe field of D1 is about 3 G for a center field of 1.3 T.  This amount of fringe field poses no problems for the RHIC beam consistent with the measurements discussed in the previous paragraph.

The magnetic field vectors in all five dipole magnets were mapped in approximately 1cm steps in three orthogonal directions, x, y and z, where x and y are defined above and z is the direction through the magnet.  The mapping was done with a sensor containing three Hall probes mounted closely at the tip of a rod, and oriented in such a way as to read all three components of the field.  These probes were moved along a long rail with a micro-step motor in the z direction. The displacement along the x coordinate was fixed by placing the long rail in holes 6.35mm (¼”) apart machined in bars mounted on the magnet shields.  Raising these horizontal bars along four threaded bars changed the elevation or y coordinate.  The y coordinate was measured with rulers having 254 m divisions.  Each magnet was mapped twice, one map at a half field and a second map at full field.  During the mapping, the pitch and roll angles of the probe were recorded as well as the temperature at the probe holder.

Corrections for small probe rotations along the x-axis and planar Hall effect contributions to the small field components were corrected for in the off line analysis by using the measured pitch and roll angles.  Once the corrections were applied an interpolation algorithm was used to produce a one cm³ grid of field values inside the magnet gap.  Figure 6 shows an example of a field map of the x component of the field in D5 at half maximum value.

The fields of D2-D5 were also modeled with TOSCA and excellent agreement between calculations and measurements were obtained.

Figure 5.  TOSCA calculations showing the fringe fields in the y-direction as a function of the distance from the center of D1.  The beam is at x=23 cm when D1 is at 2.3 degrees.

Spectrometer Detectors

Figure 6.  D5 Bx component measured 1 cm above the medium plane (y=1 cm).  The measurements were made with an average field of half the maximum value.

### Time Projection Chambers

Time Projection Chambers (TPC’s) are used as 3-dimensional tracking detectors in both the Forward Spectrometer (FS) and Mid Rapidity Spectrometer (MRS).  The TPC’s were designed to achieve an intrinsic position resolution of better than 400 µm and a two-track resolution better than 10 mm.  As the TPC’s are placed outside magnetic fields, the tracks are straight lines.  The particle multiplicities in the each TPC, including secondary charged particles, are expected to range from 4 to 120 and the particle densities between 0.01 and 0.20 particles/cm2.

Two TPC’s are located in front (T1) and the back (T2) of the D2 magnet in the Forward Spectrometer and the other two in front (TPM1) and behind (TPM2) the D5 magnet in the Mid Rapidity Spectrometer (see Figure 2).

The geometry of each of the 4 TPCs is adapted to its location.  Table 2 summarizes the dimensions and other characteristics of those individual TPCs.  T2 is built with two active volumes separated by a 5 cm spacing in a single field cage.  The two read out planes each hold 8 pad rows separated by a 12.5 mm ground bus.  Not all pad rows are instrumented in T1, T2, and TPM2, partly due to a desire to have one readout board per TPC, but also because track densities are such that sufficiently accurate measurements can be obtained with fewer instrumented rows.

The directions x, y and z used in Table 2 and throughout the paper, are defined as follows: y is in the direction of the main magnetic field (up), x is in the dispersive direction perpendicular to y and z is the length direction through the magnet, perpendicular to x and y.

Each TPC is contained in a Lucite box with two 25 µm mylar windows.  Figure 7 shows the TPM1 containment box with the field cage assembly.

In the TPC, electrons and ions are created from ionization by a charged particle passing through the gas volume.  The primary electrons drift in the electrical field towards the read-out proportional chambers, i.e. upwards in Figure 8. A minimum ionizing particle creates, on the average, ~90 electrons/cm.  A uniform drift field gradient is guaranteed by field cages made out of a series of thin aluminized mylar strips (12 mm wide, 2 mm apart) connected to a resistor chain that degrades the bottom cathode plate voltage at –5000 Volts to the ground potential near the read out plane (see Figure 8).  With a drift length of 21.8 cm, the field is 229 V/cm, yielding a drift velocity of 1.45 cm/µs.  Extensive simulations have been performed during the design of the TPC field cages using the GARFIELD code [8]. As an example, Figure 9 shows the calculated electric field lines together with the positions of the wire plane.

Figure 7. TPM1 Lucite box and field cage during construction

Table 2. TPCs main characteristics

| TPC | Dist. to Vertex (cm) | Overall (x,y,z) cm3 | Active (x,y,z) cm3 | Pad rows | Pads per row | Pad size mm2 |
| --- | --- | --- | --- | --- | --- | --- |
| T1 | 500 | 45x30x70 | 33x22x56 | 14 | 96 | 27.5x3 |
| T2 | 810 | 50x30x88 | 2x(39x22x31.5) | 2x8 | 112 | 27.5x3 |
| TPM1 | 95 | 67x34x67 | 37.5x21x36 | 12 | 96 | 29.5x3.4 |
| TPM2 | 285 | 82x30x61 | 50x22x67.5 | 20 | 144 | 24.5x4.2 |

The choice of gas is determined by the relatively small size of the magnet gaps, which entails a short drift distance and therefore favors a slow gas.  The mixture of 90% Ar and 10% CO2 gas has been chosen for its slow drift velocity and low transverse (190 µm/cm1/2) and longitudinal (240 µm/cm1/2) diffusion constants.   The TPC gas is fed to the chambers through individual mass flow meters used to adjust and regulate the gas flow. During operation, the average gas flow in each TPC is of the order of 100 ml/min and the gas exhausts to the atmosphere with an overpressure of ~1cm of oil.

A precise determination of the y coordinate requires a precise knowledge of the drift velocity in the gas volume, a velocity that must be monitored continuously.  This is achieved by a set of small drift chambers connected in series with the gas supply inlet of each TPC.  The signals induced by two collimated 90Sr sources on the bottom proportional counter are recorded by a Time to Digital Converter (TDC) and the drift velocity determined from the time difference between the two peaks in the time distribution and the distance of the two sources.

The TPC’s are read out by multiwire proportional chambers where the primary ionization electrons are multiplied through avalanche at the anode wires (sense wires) by a factor of ~4000.  Image charges are induced on an array of pads at ground potential located 4.5 mm above the sense wires.  For each track segment, the drift time provides the y coordinate, while the position and amplitude of the induced signals on the pad rows are used to extract the x- and z-coordinates. The pulse height of the induced signal on each pad, i, of an ionization cluster depends on the distance (x- xi) between the induced charge centroid, x, and the pad position, xi.  The resulting pulse height distribution, the Pad Response Function (PRF), is approximated by the Gaussian function

Figure 8.  Design of the field cage and resistor chain.  The cathode plane is at the bottom of the figure and is held at –5000 V, while the annode plane, near the top, is at ground potential.  See the text for further details.

Figure 9.  Electric field lines near the TPC read-out chamber.  All wires are perpendicular to the plane of the paper.  The positions of the pad plane and the sense (S), field (F), cathode (C) and gating (G+,G-) wires are indicated. The cathode plane (C) is denoted as the Frisch grid in Figure 8.

Table 3.  Characteristics of the grid wires.

|  | Sense | Field | Frisch | Gating |
| --- | --- | --- | --- | --- |
| Material | W (Au plated) | CuBe | CuBe | CuBe |
| Wire diameter | 13 µm | 100 µm | 50 µm | 50 µm |
| Tension | 15 g | 150 g | 100 g | 100 g |

Pi=A⋅e−(x−xi)22σx2,

where the charge distribution width x is a function of the drift length, the track angle and the transverse and longitudinal diffusion constants.  The geometry of the pad plane has been chosen so that on the average 3 to 5 pads collect the charges resulting from a primary ionization cluster.

The anode plane contains sense and adjacent field wires with a 2 mm (T1, T2) or 2.5 mm (TPM1, TPM2) pitch.  A cathode grid (Frisch grid) at ground potential is placed 4.5 mm below the anode plane and a gating grid is located 8 mm below the cathode grid.  Table 3 lists the characteristics of the various wires.

In order to have a TPC in the sensitive mode only during the drift time of beam related events, a gating grid is installed between the Frisch grid and the drift region.  In its closed state, the gating grid prevents positive ion feedback from the avalanche region near the sense wires into the drift volume.  The gating grid is opened for about 20 µs by the event trigger.  When open, all gating grid wires are at the same potential (-180V).  The closed state is obtained by applying different voltages to adjacent wires (-100V and -260V), thereby creating an electrostatic field that renders the gate opaque to the passage of charged particles.  The gating grid pulser circuit used to switch between the states follows closely the design developed for the NA49 and ALICE experiments [9].  Oscillations are avoided by adapting the output impedance to the cable.

The charge collected on the TPC pads is amplified and integrated by a low-noise, low-input impedance charge sensitive preamplifier followed by a shaping amplifier.  The pulse height and the time history of the pad signal covering the maximal drift time are fed to a switched capacitor array analog memory clocked at 10 MHz.  Those functions are achieved by the Front End Electronic cards obtained from the STAR collaboration [10].  Each card holds 2x16 preamplifier and shaping amplifier channels and stores signal in 100 ns time bins.

The gating grid pulser and the readout board are triggered by a custom built trigger/clock board as described in additional detail in the later section on trigger and readout.

The TPCs were tested with cosmic rays and the one particle resolution averaged over rows is shown in Table 4.

Drift Chambers

Table 4.  The resolution, averaged over rows, measured from track residuals.

| TPC | X(pad) m | Y(time) m |
| --- | --- | --- |
| T1 | 380 | 396 |
| T2 | 373 | 407 |
| TPM1 | 310 | 427 |
| TPM2 | 387 | 490 |

The Drift Chambers (DCs) employed in the BRAHMS detector will be described in detail elsewhere [11]; a brief overview is given below.  The DCs were designed to achieve a position resolution of less than 100 m FWHM as well as to allow the resolution of two tracks spaced more than 10 mm apart.

The forward spectrometer employs three DCs, T3, T4 and T5, placed at 10.5 m, 14.5 m and 18.4 m, respectively, from the nominal vertex.  The detectors are designed to sustain the multiplicities expected at their respective positions, i.e. from 6 to 7 for T3 and somewhat less for T4 and T5.  Therefore each module of T3 contains 10 detection planes and each module of T4 and T5 contains 8 detection planes.  Each DC contains three modules.  The sense wires are spaced 1.0 cm apart in T3 and 2.2cm apart in T4 and T5.  The active areas of the DCs are 30 x 40 cm2 for T3 and 35 x 50 cm2 for both T4 and T5

The wire directions are x, y, u, v, where the two first correspond to the x and y directions defined above, while u and v are + and – 30 degrees relative to the y direction.  Each x and y plane is repeated with a plane in the same direction but shifted half a wire distance over, to remove right-left ambiguities.  A detection plane consists of three frames made of epoxy-fiberglass.  One frame contains the alternately ordered anode and cathode wires.  Two frames with the field shaping wires are placed on either side of the sense wire frame.  The anode wire is gold-plated tungsten of 20 m diameter stretched with 50 g tension.  The cathode and field wires are made of beryllium-copper of 50 m diameter stretched with 100 g tension

The gas mixture employed was 67% Ar + 33% C4H10.  A small quantity of the ethylene is admixed at 8o C vapor temperature.

The sense wire pulses pass through a coupling capacitor and are amplified and discriminated by the custom made front-end electronics based on the ASD-8B integrated circuit developed at the University of Pennsylvania [12].  The ASD-8B is an 8-channel chip with differential inputs and outputs.  Each channel consists of a charge preamplifier, shaper and discriminator.  The boards were constructed in 16 and 8 channel versions based on the same printed circuit layout.  The front-end boards are connected to the detector with 34 pin edge connectors.  The differential output signal from the discriminator is connected to the 16-channel fast comparator board with regular twisted pair flat cables.  The ECL standard comparator outputs provide signals for the TDCs.  To avoid output to input pickup the comparators have been placed on separate boards.  The ECL output pulses were fed to the LeCroy 1877 FASTBUS TDCs on twisted pair cables and were used as the stop pulses.  The common start pulse for all the TDCs channels was obtained from the BRAHMS common start timing signal.  In order to reduce the number of TDCs channels required, a multiplexer module was constructed.  In this module each pair of signals of two 16 channel inputs is combined onto a single line placing them there at different times separated by 400 ns. The multiplexing is used for the T4 and T5 DCs.

In Figure 10 we present the deviation in position between the x and x’ planes from a prototype detector in a test run.  The figure shows that single particle position resolution is 82 m in the position resolution, somewhat better than the design figure of 100 m.

Figure 11 presents of photograph of T3 installed on the backward platform of the FS.  T3 consists of the 3 modules to the right in the picture with the electronics described above placed on top.  Immediately behind T3 follows the D3 magnet and in the background D4 may be seen.

Figure 10.  One particle position resolution of the T3 prototype detector measured during a test run.

### Time of Flight Systems

The particle identification (PID) in BRAHMS is performed in part by measuring the time of flight (TOF) of produced particles.  Since BRAHMS is a collider experiment, one must determine the location of the collision as well as define a start time in order to extract meaningful PID from the TOF information.  The start time and collision location are determined by the beam-beam array (see the subsection on beam-beam counters below).

There are three TOF “walls” in BRAHMS, one in the MRS (TOFW) and two in the FS (H1 and H2).  All three are built of a number of vertical plastic scintillator “slats” read out by two Hamamatsu R3478 photo-multiplier tubes (PMT).  One photo-tube is placed at either end of each slat.  In TOFW, acrylic light pipes with a 30 degree bend direct the light to the PMTs.  The bending direction alternates so that the tubes do not interfere and the tubes are coupled to the light pipe with 3 mm thick silicon “cookies” to allow for slight mechanical misalignments.  In H1 and H2, the photo-tubes are attached directly to the two ends of the scintillator slat with optical grease.

Figure 11.  T3 as installed in the forward spectrometer .

Two signals are extracted from the base of each PMT, one is sent to a nearby Phillips 706 or 710 discriminator used to generate the timing signal, while the other signal, after a cable delay, enters a LeCroy 1885F FastBus ADC.  The timing signals from the first discriminators are sent to the fast electronics hut, where, after a second discrimination circuit followed by a delay line, they are recorded in a Phillips 7186 CAMAC TDC.  The TDC is started by the common START and stopped by the individual signals from each PMT.  The second set of discriminators, also Phillips 706 or 710, can also be used to provide signals to produce a hardware trigger on the number of hits in a given TOF wall.

The timing information is calibrated by summing a large number of events and assuming that the earliest peak in the resulting time spectrum is due to =1 pions.  The intrinsic resolution of the individual slats has been measured in a test beam setup to be 65 ps.  The in situ detector resolution is illustrated in Figure 12 and Figure 13 with data from typical slats in H1.  Figure 12 shows the difference in time measured in the top and bottom PMT, where it is required that the track hits the slat in one of three 1 cm regions in y.  The peaks have a  of 99 ps.  Neglecting the contribution due to the finite area over which the tracks hit, this implies a time resolution of 70 ps for each measurement.  Figure 13 displays the difference between the y position of extrapolated TPC tracks and the y positions calculated from the time difference in the top and bottom PMT and shows the position resolution obtainable from a typical slat.  The broadness (RMS 1cm) is primarily due to the timing resolution.  It is noted that the position resolution is not good enough to contribute to the overall tracking, but provides unambiguous assignment of TOF hits to spectrometer tracks.

Figure 12.  Time difference between top and bottom in a typical H1 slat.  Tracks are required to intersect one of three 1cm regions along the length of a slat (y-direction).

Figure 13.  Number of tracks in H1 plotted against the difference in the y position of the extrapolated T2 track intersection with the slat, and the y position in H1 calculated by the time difference (top - bottom) for all reconstructed tracks hitting the slat

Table 5 summarizes the physical properties of the three TOF walls.  In the first RHIC run, TOFW consisted of 83 slats, but has since been expanded to 125.  The slats in TOFW are placed side-by-side along an arc with radius of curvature of 240 cm with the center located at the center of the D5 magnet, so that particles exiting D5 will have nearly normal incidence upon the individual slats.  In H1 and H2, the slats are arranged in two staggered planes, with a 1 cm gap between each slat.  This geometry allows full acceptance coverage with active scintillator, whereas placing the slats side by side would have necessitated a significant dead area due to the wrapping (Al foil and black plastic tape) of each slat.

As mentioned above, the particle identification does not depend solely on the timing resolution of the individual slats, but also depends upon the resolution of the counters used to form the common START as well as the precision with which one is able to reconstruct the position of the collision vertex and the subsequent track length through the spectrometer.  The total TOF resolution has been found to be ~100 ps in both TOFW and H1.

### Čerenkov  Detectors

C1 is a threshold Čerenkov detector installed as part of the front forward spectrometer downstream of the H1 time of flight hodoscope. It extends the particle identification obtained by H1.  The detector has the pion threshold set at 2.6 GeV/c and it can discriminate between pions and kaons up to the kaon threshold at 9.3 GeV/c. It is not possible to discriminate between kaons and protons unambiguously between the limit of time-of-flight discrimination (4 σ cut) at 5.7 GeV/c and the kaon threshold of 9.3 GeV/c.

The Čerenkov detector design was kept as simple as possible.  The radiator volume is 75 cm long, 50 cm wide and 40 cm tall and is filled with C4F10 gas at atmospheric pressure.  Čerenkov light produced by charged particles with momenta above threshold is reflected from two flat mirrors onto two photon detector planes above and below the radiation volume (see the schematic in Figure 14).  The photon detectors are closely packed arrays of 2" diameter photomultiplier (Hamamatsu H1161) arranged in two rows of 8 phototubes each.  Light reflected by the flat mirrors is funneled into the active area of the PMTs by a set of glass mirrors arranged in the shape of a truncated pyramid (Winston cones).  Figure 15 shows this array of collimating mirrors seen from the entrance window of the detector as reflected in the two large mirrors.  The number of photo-multipliers defines the segmentation to first order.  The cone of light that reaches the photon detector plane has a radius of ~6 cm and will shine on several cells.  Tracking of the charged particles is necessary to extract the necessary particle identification information.

Figure 14 Schematic side view of the Čerenkov detector C1.  For details, see the text.

Table 5.  Physical properties of the various TOF subsystems.

|  | Location | Dist. to Vertex (cm) | Number of Slats | Slat Dimensions x ´ y ´ z (cm) | Material (Bicron) |
| --- | --- | --- | --- | --- | --- |
| H1 | FFS | 860 | 40 | 1 ´ 20 ´ 1 | BC420 |
| H2 | BFS | 1890 | 32 | 1 ´ 40 ´ 1.5 | BC420 |
| TOFW | MRS | 430 | 83 (125) | 1.27 ´ 22 ´ 1.25 | BC404 |

Entrance and exit windows are made with 25.4 m (1 mil) aluminum foils mounted on the detector with O-rings.  To fill the detector with the radiator gas it was first placed in a filling vessel designed for this purpose.  The filling vessel and the detector were then evacuated simultaneously while the pressure differential on the windows was constantly monitored.  Once a good vacuum was reached air was let into the filling vessel and C4F10 into the detector.  This filling procedure was done slowly while the pressure difference between the two volumes was maintained below 33 mbar.

The average index of refraction of C4F10 for wave lengths within the photo-multiplier dynamic range was measured with beams of momenta ranging from 3 to 7 GeV/c to be (n - 1) = 138010-6, consistent with the threshold values above.

The identification of high momentum particles in the Forward spectrometer is done with the H2 time-of-flight detector and the Ring Imaging Čerenkov (RICH) detector (see Figure 3).  This detector is described in detail in reference 13.  In this section we outline the major parameters of the detector.

The RICH detector can identify pions with momenta starting at 2.5 GeV/c.  Above 26 GeV/c the ring radii of pions and kaons cannot be differentiated.  Protons can be identified well above 30 GeV/c.  The detector is situated in a low multiplicity environment of one or two tracks per event.

The RICH detector has a gas radiator volume of 878 liters filled with a mixture of C4F10 and C5F12 at 1.25 atm.  The Čerenkov light is focused onto a photon detector by a spherical mirror with a 3 m radius of curvature.

The photon detector is an array of 80 photo-multipliers manufactured by Hamamatsu Photonics (model R7600 03 M4).  Each PMT has four pixels for a total of 320 pixels.  The PMTs are mounted in four bases that contain 20 tubes each, packed as closely as possible.  The dead space between PMTs is about 2 mm.  The active area of each pixel is a square 1.2 cm on a side.  Each pixel on the photon detector has an amplifier because the PMT does not have enough gain for the FASTBUS ADC located in the Front End Electronics Hut.  The amplifiers are IO-345 designed at BNL's Instrumentation Division.

Figure 15.  A photograph of the mirror system of C1.  The picture was taken from the entrance window.  The reflection of the array of light funnels (Winston cones), that direct the light onto the PMT arrays, is clearly visible.

The spherical mirror is rotated by 8 degrees with respect to the spectrometer axis to focus the light onto the photon detector whose center is 16 degrees off the axis and well shielded from charged particles that might cause background to the Čerenkov light.  The photon detector is mounted behind a 25.4 mm (1”) thick quartz window 1.5 meters from the center of the mirror.

The index of refraction of the radiator was set at 1.001850.  Such a high index in a gas was obtained with a mixture of two perfluorocarbons: C4F10 with a measured index of n-1=1380x10-6 at one atmosphere and C5F12 with a calculated value of n-1= 170010-6 [14].  To fill the RICH, the radiator volume was first evacuated to a few millitorr, while the entrance and exit windows were protected with end caps that allowed a vacuum on both sides of the windows.  The first fluorocarbon to be introduced in the radiator was C5F12 because at normal conditions it is a liquid.  It was allowed to boil until it reached the necessary partial pressure of 630 torr.  Once that partial pressure was reached the vessel was filled with pure C4F10 to a pressure of 1.25 atm.  The index of refraction was measured immediately after filling the detector using a Young interferometer that compares the optical path in two parallel tubes, one held at vacuum and the other filled with a sample taken from the detector.  The interference fringes are detected with a PIN diode and counted with a PC after it is digitized in a Keithley’s KNM-DC32.

During the first RHIC run some data were collected with the RICH.  Even though the tracking system was not commissioned at that time it was possible to make a one to one correspondence between rings in the detector and tracks with momenta above threshold.  Figure 16 shows one event with a clear ring of at least 20 photo-electrons.

## Global Detectors

The global detectors include a multiplicity array composed of silicon strip detectors and scintillator tiles, a set of beam-beam counters and a set of zero degree calorimeters to measure neutrons.  A schematic of the arrangement of the multiplicity array and beam-beam counters in relation to other elements of the spectrometer is shown in Figure 17.

Multiplicity Array

Figure 16 An event with a ring detected in the RICH detector.

The Multiplicity Array is used to establish overall charged particle multiplicities.  An inner barrel of Si strip detectors and an outer barrel of plastic scintillator "tile" detectors measure the energy loss of particles passing through the array.  The detectors are only modestly segmented in order to minimize the complexity and cost of the array.  For a typical central collision, the segmentation results in multiple particles passing through individual elements.  The number of particles passing through each element is determined by dividing the total energy observed in that element by the corresponding predicted average energy loss for a single particle as determined through GEANT simulations.  This procedure works because there is a reasonably well defined average stopping power for the charged particles emitted in a RHIC collision.  The relationship between the measured multiplicities and the reaction centrality can be deduced based on model calculations and simulations of the array response.  A pseudo rapidity range of -2.2<<2.2 is covered for collisions occurring at the center of the array.

Figure 17 Schematic depiction of the multiplicity array and beam- beam counters relative to other parts of the spectrometer.

A view of the Multiplicity Array mounted around the beam pipe is shown in Figure 17 while Figure 18 shows a picture of the Multiplicity Array without the beam pipe.  All array elements are supported by an aluminum structure designed to have as little mass as possible and still preserve precision alignment of the detector elements.  The support splits to allow for the removal of the array from around the beam pipe without having to remove individual array elements.  The Si strip detectors are mounted close to the beam pipe in individual light-tight boxes.  These, in turn, are suspended from the Al support frame by boxes containing the front-end electronics for the detectors.  The plastic scintillator tiles are mounted directly onto the support frame.

Figure 18 Multiplicity Array with the Si strip and the scintillator tiles detectors.  The beam pipe is not seen.

During the year 2000 run period at RHIC, the inner barrel of Si strip detectors [15] consisted of 25, 4 cm by 6 cm by 300 m wafers, each segmented into 7 strips with a 0.86 cm pitch.  The detectors were located 5.3 cm from the nominal beam axis.  Both the Si strip detectors and scintillator tile detectors are arranged in hexagons about the beam axis.  For the Si strip detectors, three rows were populated with six wafers each, covering 42 bins in pseudorapidity, one row was populated with five wafers, and two additional rows were populated with one wafer each, allowing one "ring" about the beam axis to be fully populated.  The sparse population of two sides was to accommodate a clear line-of-sight for the two spectrometer arms and for reasons of cost.  The discrete component front-end electronics located in the box supporting a given wafer contains seven independent preamplifier and shaping amplifier circuits.  The amplified signals are transmitted from the Interaction Region (IR) to a remotely sited electronics room where they are processed, employing 11 bit FERA readout for the computer acquisition.

The outer detector barrel consisted of 38, 12 cm by 12 cm by 0.5 cm tiles made of BICRON BC404 [16] scintillator connected by optical fibers to adjacent photo-tubes.  The tiles were located at a distance of 13.9 cm from the nominal beam axis.  Four rows were fully populated with 8 detectors, each, one row had 4 detectors, and the last row had 2 detectors.  The partially populated rows were again to accommodate the spectrometer arms.  The tiles are independently wrapped with white bond paper and black Tedlar PVF file [17] before mounting onto the support frame.  A 2 mm deep and 1.2 mm wide groove was machined along the edge of each tile to accommodate two turns of wavelength-shifting optical fiber.  The fiber, of type K11 [18] is multiclad to increase light trapping efficiency by 50% over conventional single clad fiber, achieving an acceptance light cone opening angle of 26o and a numerical aperture NA=0.72.  The diameter of the fiber was chosen as 0.83 mm to accommodate the required bending radius of 3mm.  Both ends of the optical fiber extend for a length of 75 cm out of the scintillators and are covered with a black light tight jacket.  The phototubes used for the tile array are Hamamatsu H3178-61 assemblies.  The assembly is based on the 1.5 in. diameter tube R580-17 which has a green enhanced photo cathode.  Computer readout was done using dual-range fastbus ADCs located in the Fast Electronics Hut (FEH).

The Si strip detectors and scintillator tiles give quasi-independent measurements of the event-by-event charged particle multiplicities allowing the two measurements to be averaged in the final determination of this quantity.  The detectors also complement each other.  The signals from the plastic scintillators are very fast and suitable for use in the first level "trigger" of the BRAHMS experiment.  These detectors are also relatively inexpensive.  The Si strip detectors have the advantage of greater segmentation, allowing 42 slices in  as compared to 8 for the tiles.  The correction factor for secondary interactions, as deduced by the GEANT simulations, is also considerably less in the Si detectors, where it ranges from 6% to 30% of the total number of "hits" as compared to 20% to 40% for the scintillator tiles.  The relatively high capacitance of the Si detectors makes it difficult to develop the fast signals necessary for the 1st order trigger circuit.  In the geometry employed for the year 2000 experiment, the Si detectors subtended approximately 68% of the scintillator tile solid angle.

Figure 19 shows the individual multiplicity distributions for the Si detectors and scintillator tiles as well as the correlation between the two.  The difference in the extent of the number of hits for the two detector types reflects the different geometric acceptances.  The different end-point behavior seen for the two multiplicity distributions can be attributed to differences in the Landau distributions of energy deposited in the Si and scintillator detectors and to a somewhat larger probability of electronics saturation for the scintillators.  In general, however, a very strong correlation is seen between the two measurements.

Beam- Beam Counters

Figure 19 Particle multiplicity distributions for the Si strip detectors and scintillator tiles.  The multiplicity distribution measured by the Si strip detector is shown to the left in the figure and the corresponding scintillator tiles measurement at the bottom.  The correlation of the two measurements is shown as the scatter plot in the middle.

The beam-beam counters (BBC) are deployed to characterize collisions from a global perspective, to provide the zero level trigger, the start time for the time-of-flight measurements, and to determine the primary collision vertex position to an accuracy of the order of 1 cm.

The beam-beam counters consist of two arrays of Čerenkov detectors with one array located on each side of the interaction region. The design constraints of the system are dominated by the need to provide good start timing for the spectrometers.  The modules are designed to detect Čerenkov light from ≈1 charged particles.  The placement in a region of high pseudo-rapidity ensures that most particles have  near 1.  However this region has a high particle density, so that multiple particles striking a single module may compromise the ability to provide good timing.  This problem is controlled by the use of two kinds of detector elements: one with a small enough geometrical solid angle, so the probability for multiple hits is small even for central collisions, and one with a much larger solid angle, so the probability of receiving at least one hit is large even for peripheral collisions.

The two arrays are placed at 220 cm backwards and forwards of the nominal interaction point, employing two kinds of Čerenkov detectors in a distributed arrangement. Figure 20 shows the actual view of the two BBC arrays in the IR.  The right array has reduced azimuthal coverage because the magnet D1 moves to within a few cm of the beam pipe when placed at the most forward angle.  This reduced coverage is partially compensated by higher spatial density of modules in the array.  The angular coverage is from 1.6 to 5.8.

Figure 20.  The beam-beam counter array.  The left beam-beam counter array is shown at the top and the right array at the bottom.  The C-shape holder in the picture is part of the beam pipe support system placed ~30 cm in front of the BBC array.  The large and the small solid angle elements are clearly distinguishable in the photo.

The two kinds of modules used in the arrays. The small tube modules use a 3/4 inch Hamamatsu H5010 fast timing photo-multipliers and the big tube modules utilize the 2 inch Hamamatsu R2083 tube.  Each module has a Čerenkov lucite (UV transmitting) radiator, attached to the photo-multiplier face.  The length of the radiator is 4 cm and 3 cm for small and big tubes, respectively.  In order to ensure both good timing and single hit response, this choice was made based on test beam studies.

The Čerenkov radiator also provides some directionality selection since particles entering at large angles to the front surface or from the side and back give a much reduced signal.  The left array consists of 44 modules made up of 36 small tubes and 8 big tubes, while the right array has 35 modules in total, with 30 small tubes and 5 big tubes (see Figure 20).  Each tube has two outputs.  The signal from one output is fed to a timing discriminator situated close to the IR and the output NIM logic signal is sent on RG58 cable to the fast electronics hut, where it is regenerated and split in another discriminator.  One of the split signals is used in the trigger and the other is sent as a stop signal to P7186 TDCs.  The other output from the tube is sent, as an analog signal, directly to FASTBUS ADCs.

The prototype test was done using the secondary  beam at the AGS accelerator at Brookhaven.  A demonstration of the performance is shown as the timing between a start counter and a large tube in Figure 21.  The signals have been corrected for time walk, and the combined timing resolution of ~75 ps demonstrates that the large tubes detectors have an intrinsic resolution of ~60 ps.

The average time between a counter in the left array and one in the right array (or between a group in the left array and a group in the right) is a suitable time zero signal for the entire BRAHMS detector.  In the limit where all the beam-beam detectors are considered to be on the beam axis and all particles detected to have velocity =1, the average time of the right and left detectors is independent of the position of the primary collision vertex for any collision, since the sum of the two times equals the distance between the detector arrays divided by c, the speed of light.  Thus the time zero signal is taken to be

Figure 21 Time distribution between a start detector and one of big tube counters in a test beam set up at the AGS (see the text for further details).

T0 = (tleft + tright)/2,

where the tleft and tright are the average times for a sample of left array and right array tubes, respectively.  All times have been suitably corrected for walk (slewing correction) and all transit times through photo tubes and cables have been measured and corrected for.  The difference between left and right array times is a measure of the displacement along the beam axis of the primary collision vertex from the nominal vertex position,

z = c(tleft – tright)/2 =c((tleft –T0) – (tright-T0))/2,

where the right hand expression can be constructed from the TDC results and from the corrections for transit times.  The overall resulting uncertainty in the determination of the vertex from the beam-beam counters is better than 1.2 cm.

The tubes also exhibit good multi-hit resolution.  Figure 22 shows a normalized ADC spectrum from minimum bias Au-Au collisions at 130 GeV for one of the 2” big modules.  Hits up to ~5 are clearly visible in the spectrum.  It has been established by calibrations with a laser in a laboratory setup that the big modules have a linear response for the range of interest (up to 15 hits).  Examples of the use of the BBC arrays for the spectrometers (T0 and z determination) and as a multiplicity detector are shown in the “Spectrometer Performance” section.

Figure 22 Normalized beam-beam counter ADC spectrum for minimum bias Au-Au collisions obtained with a 2 inch big tube.

### ZDCs

The Zero Degree Calorimeters are designed to measure the luminosity of the colliding beams.  This capability allows the beam operators to tune the machine [4] and provides a means to compare results between different experiments. In addition the time difference between the two ZDCs may be used to measure the position of the primary interaction vertex and to deduce the impact parameter of the event[6].  The BRAHMS experiment also uses the ZDCs to study the heating of nuclei by photons in extremely peripheral collisions [6].  Finally the pattern of energy and time signals from the individual modules in each ZDC allows the rejection of certain types of background events a feature that will become increasingly important as the luminosity of RHIC increases.

High energy collisions of nuclei usually lead to the emission by evaporation of spectator neutrons from both nuclei.  At RHIC energies evaporation neutrons diverge by less than 2 mr from the beam axis.  It is these neutrons which are detected by the zero degree calorimeters (ZDCs)[4].  In this “zero degree” region produced particles and other secondary particles deposit negligible energy when compared with that of beam fragmentation neutrons.  The coincidence of both ZDCs serves as our minimum bias trigger and luminosity monitor[5].  The neutron multiplicity is also correlated with event geometry [6] and may be used to measure collision centrality.  Since they are the only common detectors between the four RHIC experiments, cuts on the ZDC response provide a way to compare results among the RHIC experiments.

Each calorimeter is positioned at a distance of 18m from the nominal interaction point.  Each calorimeter is divided into three modules, which is mechanically convenient and also allows a measurement of the energy deposition along the beam axis.  This is a useful tool to discriminate against background particles entering the calorimeter from the back end.  Since space is very limited the ZDC use a very dense absorber, tungsten, and use optical fibers to sample only the high energy particles that produce Čerenkov light.

The difference in flight time from the vertex to each ZDC can be used to measure the vertex position while the average of the two times might be used as a time zero signal.  After correction for the dependence of time on signal amplitude (time walk) a resolution of 170 ps is obtained.  The resolution is dominated by the difference in path lengths of light from neutrons hitting the top or bottom of the ZDC.  The timing resolution of both of the ZDCs leads to a ZDC vertex resolution is 3.6 cm.  In general the BBC array provides better time resolution and is used as the vertex position and time zero signal (see the previous section).  In 2000 the distribution of vertices was approximately Gaussian with a width of 70 cm

For particles hitting the ZDCs from the front the first module should fire about 1.5ns before the third module.  For particles that hit the ZDCs from the back the time difference is reversed.  In BRAHMS this effect is used to reject background events caused by beam-gas collisions.  The magnitude of the effect depends sensitively on the quality of the vacuum in the beam pipes in the BRAHMS interaction region.  Typically only a small fraction of the energy in each event is deposited in the third module.  Thus a large energy deposit in the third module is a sign a particle may have hit the ZDC from the back.

From reactor tests at Texas A&M University it was concluded that the optical fibers could tolerate a dose of more than 500 krad before needing to be replaced [4].  During the first year of RHIC running the ZDCs showed no radiation damage.  It is planned to measure the radiation dose with radiochemical monitors placed on the ZDC and to monitor for radiation damage by the position and width of the one neutron peak.

The ZDCs are also essential for the study of very peripheral events in which the nuclei do not touch but interact via an exchange of photons [6]. Such events are selected by requiring that there are no hits in the beam-beam counters and the multiplicity array.  In such events the energy absorbed by the nuclei causes neutrons (and some protons) to be emitted with very little transverse momentum.  These neutrons can then be measured in the ZDCs.  Since each nucleus interacts independently we expect no correlation between the number of neutrons in the left and right ZDCs, a feature that is observed in the data. However for nuclear events the number of neutrons seen in each ZDC is correlated since both nuclei have the same number of wounded nucleons.

# Trigger and Data Acquisition

At full luminosity the Au-Au collision rate is ~1.2kHz of which about 10% are characterized as the most interesting central events with the highest number of tracks in the spectrometers. The distribution of collisions along the beam axis is expected to be Gaussian with a RMS of ~18cm at the beginning of a fill and growing to ~25 cm over its lifetime of ~8 hours with the storage RF rebunching implemented.  The spectrometers are designed to accept particles with optimal acceptance when the vertex is within 15 cm of the nominal IP.  The rate of central collisions within this acceptance is thus 60 Hz.  The design criteria for triggering, readout electronics and DAQ capabilities is to be able to collect these central events, and a scaled down sample of other event classes with an overall data taking rate of ~100 Hz.

## Trigger implementation

From a data acquisition point of view, the experiment consists of 12 detector systems that compromise about 500 channels of PMT, for which ADC as well as high resolution TDC information are extracted, about 450 channels of PMT for which only the ADC is read out, ~2000 channels of timing information from the drift-chambers (DC), and 160 ADC channels from the silicon array.  By far the largest amount of data comes from the TPCs, which record up to 158 time bins of 10 bit ADC data from up to 1152 pads in each TPC.  Apart from the customized readout of the TPC’s, the experiment utilizes conventional CAMAC and Fastbus modules for the digitization.  It was also decided to implement the trigger system using NIM and CAMAC electronics, in part for reasons of cost, but also because all the readout systems impose fixed dead-time, and the collision rates are modest.  The design and timing were also driven by the decision to place most of the trigger and readout electronics in a trailer (FEH), placed just outside the shielding wall in the IR area.

There are two levels of triggers, denoted level 0 (LVL0) and level 1 (LVL1).  LVL0 is defined as those triggers that gate the ADCs, TDCs, starts the TPC gating grid and the readout cycle of the TPC FEE cards.  LVL1 is a trigger which subsequently interrupts the VME processors that handle the readout of CAMAC and Fastbus modules.  For the 2000 running only LVL0 triggers were implemented.  The triggers defining collisions come from two primary detector systems namely the ZDC and the BBC.  They are:

- A left-right coincidence of the ZDC timing signals with a narrow time overlap of ~5 ns and with a summed energy signal from each module with energy > 25 GeV in coincidence with the RHIC RF clock signal.

- A coincidence between the left and right arrays of the BBC with a requirement of 2 or more hits in a subset of Big Modules.  A hit is defined as any analog signal that fires the discriminators set roughly at 40% of the signal recorded by one relativistic charged particle.  The hits are taken from the big tubes (see the BBC section) at present, but the electronics allows for the selection of any subset of tubes to be involved in the decision.

The first trigger is the most inclusive selecting not only all nuclear reactions, but also the electromagnetic dissociation peripheral collisions with the cross section of about 11 b. The second trigger selects about 60% of the nuclear collision events.  In the first RHIC run these two triggers were sufficient to collect all data of interest, since the machine luminosity only reached 10% of the nominal value.  Thus all LVL0 triggers also acted as LVL1 triggers.  At a later stage a centrality LVL0 trigger will be implemented to select the 10% most central collisions.  This will be done by requiring collisions by means of the basic BBC trigger, and then further require a large deposited energy in the Tile Multiplicity Array together with a cut on the ZDC energy.  Based on studies of existing data already such a trigger will select about 20% of all nuclear events, and all of the 10% most central events with an overall rate of about 200 events/s at full luminosity.  In addition to the physics triggers the setup allows for various pulser events interspersed in the data stream.  A rough vertex selection based on a making a cut on the timing difference between the left and right ZDCs (or BBC) counters will be implemented at LVL1 using a time-to-amplitude converter and rejecting collisions with vertices outside a narrow range of 15 cm.

Table 6 Trigger Timing

|  | Time (ns) |
| --- | --- |
| Collision | 0 |
| BBC output | 35 |
| BBC at FEH | 150 |
| LVL0 input | 320 |
| LVL0 out | 440 |
| ADC, TDC gates | 490 |
| TPC readout start | 1030 |
| TPC digitization start | 12000 |
| TPC transfer complete | 1500000 |

The physical layout of the experiment, the trigger electronics and readout electronics result in a timing sequence with characteristics as shown in Table 6.

### TPC readout

From each TPC up to 1152 pads each with 158 time bins are read out.  The START FEE cards digitize 512 time buckets, but the readout board was reprogrammed to only read and transmit the first 158 buckets sufficient to cover the drift time of 13 s in the four TPCs.  The FEE-cards connected on the detector each read 32 pads and is connected to a readout board.  The FEE cards and readout board are electronics from the STAR experiment, procured as part of a collaborative agreement.  The readout-board is triggered from a trigger/clock module, and the digitized data is transferred from the board on a Giga-byte fiber Link connecting to a VME readout board.  For a description of the FEE and readout board see ref. [10].  The two other components are described in the following sub-sections.

### Trigger Clock module

The BRAHMS data acquisition (DAQ) is event driven and the TPC readout is driven through a special Trigger/Clock module that starts the readout cycle on the readout boards as well as supplies the clock signals for the Switched Capacitor Arrays.  The trigger-clock module accepts up to 3 different external NIM inputs, corresponding to different trigger types. The input starts a sequence of reading a command word from a register, synchronizing the SCA clock to the command clock (RHIC 9.2 MHz RF clock), sending the appropriate command to the readout board and thus initiating the readout and digitization of the FEE cards.  An output synchronized to the SCA clock is sent to the TPCs and opens the gating grid.  The synchronization is essential in reducing the RMS noise.

### VME receiver module

The BRAHMS TPC receiver board is a VME based board that performs initial pedestal subtraction, pedestal suppression, and sorting of the TPC data.  Each receiver board can accommodate the data from one TPC.  One receiver board is dedicated to each TPC and is interfaced to the readout board through a 1.5 GB/s fiber optical link.  Data is transferred from the readout board to the receiver board at a rate of 20 bits every 34 ns using Hewlett Packard GLINK transmitter and receivers.

The ADC conversion time dictates the maximum DAQ rate in the first step.  For 158 time bins this amounts to 1.2 ms.  Additional dead times are introduced when transferring data from the receiver board to VME processors, an amount depending on the filling factors of the TPC.  The charge from each time bin is digitized in a 10 bit ADC’s on the FEE cards.  An important criterion was to make processing on the receiver board simple but fast, and postpone more sophisticated computations to the VME processors, that are programmable in C++.  A dedicated VME processor board is interfaced to each receiver board, and performs reformatting of the data.  The receiver board is therefore required to perform a "simple" pedestal subtraction, data ordering, pedestal suppression and clustering definition using CLPI chips and Field Programmable Gate Arrays (FPGA).  The cluster algorithm is designed to keep two time bins above and below clusters, a cluster being defined as a sequence of 2 or more time bins with ADC values above their pedestal value.  The data flow on the receiver board is shown in schematic form in Figure 23.

The tasks performed by the BRAHMS receiver board can be divided into a 3-stage pipeline:

Event Readout.                     .
Readout of the ADC data over the link: Data is read over the link at the rate of 20 bits of data and 4 bits of control information every 34 ns by the optical receiver.  The 20 bits of data corresponds to ADC values of 10 bits from two time pads.  The control bits indicate if a transfer carries data or a special word.  Consecutive data read over the link correspond to data from different pads with the pattern repeating itself for every 1152 pads for subsequent time bins.

Figure 23.  Schematic diagram for the TPC VME receiver board.  The components described further in the text are the pedestal and data SRAM, the FPGA and the FIFO.

- Data Sorting and Pedestal Subtraction. 
At this stage pedestal values are subtracted from the data, which are stored into static memory (SRAM) with 10 bits of data and two control bits indicating carry and overflow.  The pedestal values and the addresses of where to store the ADC value from a given pad and time bin are kept in two r/w SRAM tables.  It is noted that the pedestal value subtracted is the true pedestal for the ADC distribution for a given bin plus a threshold value.  The threshold is added back to the subtracted ADC value in the VME processor.  The processing is done at a rate of two words per 34 ns to keep up with the data being received over the link.  It is done in 4 processing units, with one 10 bit word being fed into each unit once every ~68 ns.                           . 
Upon the completion of event readout the data that reside in the 4 SRAM data banks are pedestal subtracted, sorted and consecutive memory locations hold data in time from the same pad.  Each of these data banks consists of 256k of 12-bit memory.  Two sets of the SRAM data banks exist, and alternate banks are used from one event to the next.

Pedestal suppression and transfer to FIFO.
Once the data has been buffered and if the previous event has been totally transferred to the receiver board the pedestal suppression stage starts.  The receiver board has two complete sets of SRAM data banks allowing for suppression and transfer of one event while being ready to receive the next event. This stage consists of determining if the data are part of a "simple" cluster, tagging the beginning of the cluster and storing the data in FIFOs.  Once a cluster is detected, the bins that constitute the cluster are written to a 8k 16-bit FIFO together with a header and the data from the time bin preceding and the bin following the sequence.  The FIFO is being emptied by reads from the dedicated VME PPC2505 processor over a dedicated VME back plane using VME32 addressing block mode.  The transfer rate is the limiting factor in the overall data taking rate.  For pedestal data where all time bins are readout the limit is ~10 events/s, while for typical data the limit is ~ 150 events/s.

Figure 24 BRAHMS DAQ data flow

The VME processor has 64 Mb of local memory enough to keep a large number of events.  The main task of the processor is to reformat the time ordered data.  This reformatting is relatively simple, consisting in adding the threshold value back to the ADC values, and adding in each data sequence the row number, the pad number and the start time bin.

## Data Acquisition Overview

The BRAHMS Data Acquisition is based on a distributed system of loosely coupled front-end processors that are being interrupted by actions of the trigger systems, and a high-end workstation that serves the functions of event builder and user run-control.  The components and the performance are briefly described below. The architecture is illustrated by the flow of data in Figure 24.  The front-end VME processors initiate and interrupt the readout of data from the crates.  The event fragments are buffered in local memory and transferred via a100 MB Ethernet to a high-end Unix server, and the event builder program assembles a complete event.  The server has a disk spool area capable of holding up to 100 Gb worth of data and acts as a buffer for the raw data files before they are sent, using the pftp protocol, via a Giga-bit link to the HPSS storage at the RHIC Computing Facility (RCF).  The files on the spool area are also available via NFS mounting for rapid analysis from the workstations in the counting house.  The DAQ design is fairly typical for a modern system, with the exception of the method employed for maintaining synchronization.  This method is unique and is described below in some detail.

### Hardware

All of the analog signals from PMTs are delayed via RG58 cables and fed to LeCroy 1885F 96 channel Fastbus ADCs.  The crate is controlled by a triple width STRUCK Fastbus Interface (STR-340 SFI).  The Fastbus interface has up to 4 interrupts of which two are used by the trigger, one for physics events and one for synchronization events.  The VME processor is situated in a 2 slot dedicated VME back plane.  A Fastbus arrangement is used for the LeCroy 1877S pipeline TDCs that read out the drift chambers.  The Fastbus crate housing the DC readout is situated on the spectrometer platform.

The high-resolution timing signals from the TOF detectors, and the beam-beam counters are read out by P7186 TDC CAMAC modules. Altogether 6 CAMAC crates in the experiment have Kinetics System 2132 Crate controllers connected to the SBE VME processors [19], which in turn are controlled by another MVME 2305 or MVME 2307.  The VME processors all run under the VXWorks embedded operating system.

The Unix server is a 6 processor SUN E3500 running Solaris 8 and has attached 100 Gb of disk with the main function of serving as a spool buffer for the raw data before these are sent to RCF.  The network connection used for sending raw data to RCF is a dedicated GB fiber network utilizing jumbo frame to allow for maximum data rate.

### Software

The DAQ software is also distributed across the processors.  The time-critical software all sits on the front end VME processors, the time-critical synchronization is done via the trigger logic by interrupts for the beginning of an event, and by setting a ready level when completed.  The next level of software runs on the Unix server.  This consists of two main components, namely the event builder (i/o-server) and the runcontrol program that is responsible for all command communication between the distributed processors, and the i/o server and the user control interface.  Each of these components is described in some detail below.

### Event Builder components (i/o- server)

The event builder has several major components, most of them implemented as threads inside the event builder process.  The receiver checks data validity, creates the event fragments and passes them to the event assembler, which matches them into complete events.  A data mapper receives assembled events and creates the mapped data records.  This process does not perform any re-mapping of the data.  An output manager takes and distributes them to output data streams on disk, net connections and possible local tape

### Event synchronization

An important issue in a loosely coupled distributed system is to ensure data-integrity and synchronization.  This is achieved in the present system, by having the trigger system set up synchronization events that causes the front-end processors to tag this sub-event as having a different data type.  These are presently set at a rate of 1 Hz, with typical 100 physics events between each occurrence.  The synchronization trigger causes a separate interrupt of the VME processors.  In the case of the TPC readout this is indicated by the ‘DAQ’ word sent to the TPC readout board from the trigger/clock module.  If the event builder sees a mismatch, i.e. the data fragments do not all correspond to such a synchronization trigger, it will reset the data taking via a command to the run-server, and will discard all events from the previous accepted synchronization event.  Such failures are rarely observed.  When they do occur it is typically due to loss of a single front end node e.g. from a power failure or a startup after configuration changes have taken place.

### Run Server

The main purpose of a separate run server program is to provide a well-defined interface between the logic part of run control, implemented as a state machine, and the user part of run control.  The run server interface is defined by an IDL interface that contains methods to initiate run state transitions, such as "begin run" or "end run", and methods to query the run server for information, such as run state, run number, etc.  By calling the control methods defined in the Run Server interface a run control program can cause state transitions in the run server.  During a state transition, the state machine issues commands to the front-end nodes and the event builder causing them to start or stop taking data.  In addition to reacting to user interface commands, the Run Server has a background task to poll the front-end nodes and the event builder for information.  The background task performs check on the collected data and if necessary, can force state transitions or initiate other actions and change the run state.

The user interface to the run control is a graphics interface implemented in Java.  This interface is running on local Linux workstations in the counting house.  In addition to the normal tasks of defining configurations for run, starting and ending runs, it gives the user access to details of events, front-end node-activities, log files etc. Configuration changes, and run conditions are recorded in a local MySQL database.

### Performance

The DAQ system, for Fastbus and Camac readout alone, is capable of handling up to 1000 events/s, equivalent to 20 Mb/s.  For real events the limits are about 500 events/s as set by the TPC readout board.  It may be further reduced depending on the amount of data after pedestal suppression on the TPC receiver board.

During the final weeks of the 2000 run the collision rate at the BRAHMS IR was about 200 Hz.  The data taking rate was ~100 Hz with an average event size of 60 kb and dead times of ~60-70%.  The event size was higher than anticipated, mainly due to a large number of clusters in the TPM1 the one nearest the IP.  The amount of data recorded in this fairly short run was about 2 Tb.

# Online Monitoring

Online monitoring in BRAHMS was accomplished in the 2000 RHIC run using a variety of methods and software.  One piece of software is a monitoring program which samples data from the event builder and increments the various diagnostics spectra generated by the selected detector groups.  The generation of a ROOT tree [20] from a subset of data, right after a particular run was completed, was found to provide a rather powerful near real-time analysis method.  The TPC tracking performance was monitored separately by sampling the TPC data and plotting various projections of the pad hits and time distributions.

## Monitoring Program

The program to monitor BRAHMS data online has three components.  The components consist of monitor modules, a “super-module”, which acts as a manager of the monitor modules, and a main program which creates the “super-module” and adds all of the monitor modules depending on which switches the user sets at startup.

The primary component is the monitoring module, designed to work standalone in ROOT or to be added to a “super-module”.  The monitoring module as presented to the operator has two lists that can be clicked on to display either a single histogram or a set of histograms which we call a picture.  In addition to these lists there are buttons to control input of the data from the event builder or from a raw data file as well as to clear arrays or to integrate the contents of one dimensional or two dimensional histograms.

The component used most often in the experiment is a “super-module” which has the capability of loading a number of the monitoring modules and keep track of them in a list.  It also has a number of buttons to control event input from the event builder or raw data file as well as a number of “global” operations such as clearing all arrays of loaded modules.  Once a particular detector module is loaded, its array incrementing routines are executed event by event.  The arrays and pictures for each detector module are presented to the operator as tabs in the main operator console.  Once a detector module is selected, the list of defined arrays and pictures for that module are presented to the user.  In addition, buttons for each monitor become visible which operate only on the data for that particular monitor.

The third component is a program that can be executed by the user to create the “super-module” and add the different detector modules according to switches set in the command argument.  There are typically several of these programs with different combinations of detector monitoring running on different nodes.

## Tree Program

A program was written to transform the raw data into some simple structures, useful for monitoring and diagnostics, structures that were written in a ROOT Tree and therefore could be read back and used for analysis.  This added analysis capability was found quite powerful, in that spectra with different gates could be generated very quickly as the need arose.

The program is similar in spirit to the monitoring program.  There is a base class which handles all of the complex ROOT features and inherited classes that perform the detector specific work to extract the proper data from the raw data stream and transform it into the ROOT Tree structures.  A main program creates all the elements and adds them to a list and then executes the data generating routine on an event by event basis.

## 3D Monitoring of TPCs

It is very important to have a graphics view of the performance of a TPC.  Many problems were diagnosed by looking at track patterns from various samples of events.  A program was developed which read in the raw data, extracted the TPC information and made pictures of the pad hits.  Both projections of the pad hits as well as 3D representations of the tracks were generated.  Figure 25 shows an example of the TPC viewer used during the run.  The top and bottom left plots show the y-z, x-z and x-y projections, respectively.  The bottom right plot shows a 3D representation of the track through the TPC.  The 3D representations of the tracks could be rotated to achieve better views of the tracks.

In addition, clicking on a pad in one of the projected views would generate 1D histograms of time, pad or row depending on the selected projection.

Figure 25 Online event viewed with the TPC 3d viewer (see the text for further explanation).

# Spectrometer Performance

Any collider experiment presents the challenge of determining where a particular collision between a particle from one beam with one from the other beam occurs.  The precise determination of the collision vertex is important for several reasons.  First of all, the distance a particle travels to the respective detectors depend on the position of the collision vertex and therefore is essential for particle identification by time of flight.  In addition, the acceptances of the spectrometers vary with the position of the collision vertex as does the determination of the polar angle of the particle momentum.

As noted above, BRAHMS has three detector systems with vertex position determination capabilities, the beam-beam counters, the ZDCs and the MRS TPM1.  The BBC array and ZDCs provide a determination of the vertex position by a measurement of the difference in arrival times for particles that hit the left and the right hand detectors of one of the systems, while the vertex determination from the TPM1 is achieved by projecting the found tracks to the beam pipe.

Figure 26 shows the correlation of the vertex determined from TPC tracking with the one from ZDC timing.  An excellent correlation between the two determinations is apparent.

The multiplicity array combined with the beam-beam counters cover the pseudo-rapidity range from -4.3 to +4.3 and may thus provide a rather complete measurement of the angular distribution of charged particles emitted from a heavy ion collision.  The three detector types involved do not measure number of particles directly: the Si strip detectors and the scintillator tiles measure energy loss, while the beam-beam detectors measure Čerenkov light.  By appropriate normalization to the average energy loss in each detector element and careful simulations, the signals from the multiplicity array can be translated into number of hits on a statistical basis.  The light spectra from the beam-beam counters, as shown in Figure 22, carry the needed multi hit information to also yield the received number of hits.  An example of dn/d distributions obtained from peripheral and central Au+Au collision at a center of mass energy of 130 AGeV is demonstrated in Figure 27, taken from ref. [21].

Figure 26 Correlation between vertex determination from the ZDCs and the TPM1 time projection chamber.

Particle identification is achieved by the use of the momentum measurement obtained with the tracking through two TPCs with a dipole magnet in between and the data from the time of flight arrays located behind the tracking detectors.  As noted above, the particle identification is dependent upon an accurate determination of the collision vertex.

Figure 27 Pseudo-rapidity distributions from 5% central Au+Au collisions at 130 GeV/u center of mass energy measured with the multiplicity array, beam-beam counters and TPM1.  The results from the tiles (open squares) are systematically higher than the Si strip data (open circles) in the mid-rapidity range, a feature that is not understood quantitatively.  The results from TPM1 (stars) agree quite nicely with the tile and Si at the same .

Figure 28 presents an example of the measured Time-of-Flight versus momentum of positively charged particles from Au+Au reactions at 130 GeV/u obtained in the MRS in the 2000 run.  Pions and kaons are clearly resolved up to 1.5 GeV/c and kaons and protons beyond 2.4 GeV/c.

For H1 in the forward spectrometer the situation is slightly more difficult since the particle momenta are much larger and the path length more dependent upon the vertex position than is the case for the MRS.  The FS was operated at an angle of 4, and the contribution of the uncertainty in the vertex position z to the path length is l= z/cosθ.  At an angle of 4 an uncertainty in the vertex position of 1 cm corresponds to a timing uncertainty of 30 ps for particles with =1.  A TOF versus momentum plot for positively charged particles is shown in Figure 30.  The reaction angle is 4 and the reaction is Au+Au at 130 GeV/u.  Pion and kaon separation is achieved in a narrow momentum interval below 2.5 GeV/c, while the protons are separated from  and K up to 4.5 GeV/c.  Even if this is as good (or better) than the separation in the MRS, the result calls for the additional information from the C1 Čerenkov counter, because the momentum spectrum is much harder at 4 than at 90.  The C1 Čerenkov counter has been successfully incorporated in the 2001 run, but was not available in 2000.  The measurements used for Figure 28 and Figure 30 are described in detail in ref. [22].

Figure 28 Time of flight of positively charged particles in TOFW as a function of reconstructed particle momentum obtained from tracking in TPM1 and TPM2 through the D5 magnet.

With the full implementation of the backward portion of the Forward Spectrometer in the 2001 RHIC run, we have begun to utilize the RICH detector.  Figure 29 shows the particle identification obtained in preliminary analysis of data collected with the RICH detector.  This plot was generated projecting tracks from the drift chambers in the BFS into the RICH and extracting radii from PMT hits in the RICH phototubes.  We observe a very nice separation of π/K to momentum beyond 15 GeV/c.  Low statistics prevent us from observing a definite proton line, but the preliminary result is that the detector is working very well.

Figure 30 Time-of-Flight versus track momentum for positively charged particles.  The TOF was measured by H1.  The dashed lines indicate expected particle TOF  2 sigma for pions, kaons and protons.  The solid curve indicates the expected deuteron time of flight curve.

Figure 29 Preliminary PID obtained from the RICH detector.  The lines indicate the theoretical positions of the particles in the plot.

# Acknowledgments

The very professional efforts of the RHIC team to commission the accelerator and deliver the beams are gratefully acknowledged.  This work was supported by the United States Department of Energy grants DE-FE05-86ER40256, DE-AC02-98CH10886 and DE-FG03-93ER40773, by the Danish Natural Science Research Council, by the Norwegian Natural Science Research Council and by the Polish ministry of Scientific Research.

# References

- T. Abbott, et al., Nucl.. Inst. Meth. A290 (1990) 41.

- L. Ahle et al., Nucl. Phys. A590, (1995) 249c; K Shigaki et al., Nucl Phys. A590 (1995) 519c.

- GEANT, Detector description and simulation tool, CERN Program Library Long Writeup W5013, CERN, Geneva (1993); http://wwwinfo.cern.ch/asd/geant.

- C. Adler et al., Nucl.. Inst. Meth. A470 (2001) 488.

- A. J. Baltz, C. Chasman, and S. N. White, Nucl.. Inst. Meth.  A417 (1998) 1.

- H. Appelshauser et al., Eur. Phys. J. A2 (1998) 383.

- TOSCA reference manual, Vector Field Lmited, Oxford, England (1992).

- R. Veenhof, The Garfield Program v.7.01, http://consult.cern.ch/writeup/garfield

- S. Afanasiev et al., Nucl. Inst. Meth. A430 (1999) 210; ALICE Collaboration, CERN/LHCC 2000-001 ALICE-TDR-7.

- S. R. Klein et al., IEEE Trans. Nucl. Sci., 43 (1996) 1768.

- In preparation.

- F.M. Newcomer, S. Tedja, R. Van Berg, J. Van der Spiegel, H.H. Williams, IEEE Trans. Nucl. Sci. 40 (1993) 630.

- R. Debbe et al., Nucl. Inst. Meth. A371 (1996) 327.

- T. Ypsilantis, Proc. of the Symposium on Particle Identification at High Luminosity Hadron Colliders, T. J. Gourlay and J. G. Mofin, ed. Fermilab, (1989) p. 133.

- Micron Semiconductor Inc, 126 Baywood Avenue, Longwood, FL, USA.

- BICRON, 6801 Cochran Road, Solon, Ohio, USA.

- DuPont, P. O. Box 88, Sheridan Drive and River Road, Buffalo, NY 14207-0088.

- Kuraray America, Inc., 200 Park Avenue, New York, NY 10166.

- Kinetic Systems, Inc., 20 Arboretum Rd., Boston, MA 02131.

- Rene Brun and Fons Rademakers, Nucl. Inst. & Meth. A 389 (1997) 81.

- I.G. Bearden et al., the BRAHMS collaboration, Phys. Lett. B, submitted for publication.

- I.G. Bearden et al., the BRAHMS collaboration, Phys. Rev. Lett. 87 (2001) 112305-1.

Figure comments:

Fig. 1 is from CDR and not quite up to date. And the ordinate should be marked pt .(old 6)

Fig. 2. A “DX” should be added to the second beam magnet and the first should be “DX” rather than “Dx”.  In general “labels” should be moved out in the clear space if possible at all. The mid rapidity TPCs should be “TPM1” and “TPM2” (that’s a spokesman decision!). The color green (or light green) used for arrows etc. should be avoided as it does not print in grey tone. Use blue or black. If possible, it would be nice to add “Nominal Vertex”, but the figure is quite full there. (old 1)

Fig. 3.It prints OK in grey tone on a good printer. I chose to show FS because it is the most distinctive feature of the BRAHMS detector. (old 2)

Old 3 is gone, it is repeated in old 19, and appears then as new 11.

Old 4 is gone, looks too much as 802 or 866.

Fig.4. Prints ok in grey tone on a decent printer. (old 5)

Old fig. 7 I have dropped;  I could not understand it. New figure 5 will suffice for fringes.

I also cruelly killed old figs.9 and 11, because the point is not an evaluation of TOSCA, but rather  to describe the methods we used in the design and the results of the constructed magnets.

Old 12=new 7, which prints ok in grey.

Old 13=new 8. –5000V has to be shown at the bottom, and 0 V at Frish and pad plane. Two “wires” below gating grid should be deleted.

Old 14=new 9, exists in version that prints better in grey.

Old 15 and 16 are gone, I did not think they enhanced the paper.

Tables 2,3 and 4 in different ordering, and the old TPM1 totally deleted.

Old 17 killed, and the DC section made ultra short awaiting a separate paper.

Old 18=new 10, old 19=new 11, old 20=new 12.

Old 21 has a new version (not inserted) as new 13.

New 13 is an update of old 21. New 14 = old 22. New 15 = old 23.

Old 24 dropped, new 3 does the job.

New 16 = old 25. New 17 = old 26. New 18 = old 27-2. New 19 = old 28. Old figure under 28, not numbered is deleted. New 20 = old 29.

New 21 = old 30. New 22 = old 31 – that’s at least the intention. Old 32 is out, it mainly shows how badly calibrated we were at the time.

New 23 = old 37. New 24 = old 38. Old 39 is out, it is not very exciting and not needed.

New 25 = old 40.  New 26 = old 33.

Old 34 is out, it will be done much better as we learn.

New 27 = one of those to be thrown out, dn/deta, maybe edited.

New 28 = old 35. New 29 = old 36, but hopefully in a new and better edition by Djamel.

Table comments:

Table 1 needs the Bdl values.

Table 6 killed – Čerenkov RESULTS would have been a nice complement, but we don’t have them.

Photo-tube comment: will somebody with insight check and correct Hamamatsu numbers?

Table 7.  Physical properties of the various TOF subsystems.

|  | Location | Dist. to Vertex (cm) | Number of Slats | Slat Dimensions x ´ y ´ z (cm) | Material (Bicron) |
| --- | --- | --- | --- | --- | --- |
| H1 | FFS | 860 | 40 | 1 ´ 20 ´ 1 | BC420 |
| H2 | BFS | 1890 | 32 | 1 ´ 40 ´ 1.5 | BC420 |
| TOFW | MRS | 430 | 83 (125) | 1.27 ´ 22 ´ 1.25 | BC404 |

Both the design of the STAR readout board as well as the BRAHMS receiver board allows for twice this rate. When operating at that speed though the readout board experienced race conditions, causing additional noise in the data.