## Slide 1: Status of the PHOBOS DAQ
How we are with the plans
DAQ Bottlenecks, results from recent test runs
Problems remained
Potential improvement of DAQ components
How to increase the disk throughput
Things to do

## Slide 2: Schedule

## Slide 3: DAQ Bottlenecks
Silicon to Mercury. 700 ev/s, 200 MB/s, based on single event data taking time = 1.4 ms
Mercury to UltraSPARC. (hard 80 MB/s, soft 64 MB/s, achievable 50 MB/s) DAQ instantaneous: 520 ev/s, 52 MB/s, 40% CPU DAQ sustained: 330 ev/s, 33 MB/s
UltraSPARC to disk. 450 ev/s, 45 MB/s, 4% CPU if record = event 35 MB/s
Mercury to disk. 280 ev/s, 28 MB/s, 50% CPU DAQ program. There is a room for improvement.
Disk to phobosx: 280 ev/s, 28 MB/s
Simultaneous reading and writing: 40 MB/s write (4%CPU), 28 MB/s read (34%CPU)
Disk to RCF. 250 ev/s, 25-30 MB/s using test program.
FASTBUS + VME combined bandwidth 6 MB/s
No FASTBUS, real data from Silicon to Mercury, simulated data in the Mercury.

## Slide 4: Problems
Wrong pedestals used for compression. Compression factor = 1, expected 3-4
FASTBUS data transfer over Ethernet failed during last tests
Fast transfer of FASTBUS and VME tested standalone but not yet in the DAQ.
Recorded data files have final format but bad compression and missing FASTBUS.

## Slide 5: Pipelining
Fibers from DMU crates @ 2*120 MB/s
Workers
22*15 MB/s
Collector
FPDP link to UltraSPARC
@ 50 MB/s
Should work up to 700 Hz

## Slide 6: Pipelining, potential improvement
Fibers from DMU crates @ 120 MB/s
Workers
22*15 MB/s
Collector
FPDP link to UltraSPARC
@ 50 MB/s

## Slide 7: Increasing the writing performance
Additional hardware:
FPDP-PMC interface, $2k
ROUT PMC board, $2k
Dual SCSI-PCI interface, $1k
PCI-Gigabit interface, $1k
Disk enclosure, $2.5k
4*180 GB disks, $8k
Advantages:
Known parts.
Increased disk space
Disadvantages:
Ultra 5 will be overloaded. Upgrading to double CPU will cost $15k
Gradual: Using two UltraSPARCs and two disk arrays

## Slide 8: Increasing the writing performance
Additional hardware:
Fiber channel - RACEway, $6k?
Fiber channel RAID controller, $4.4k
Fiber channel HBA, $2k?
PCI-Gigabit interface, $1k?
Advantages:
Mercury to disk directly @100 MB/s
Disadvantages:
No file systems on the disk array. How PFTP to RCF?
All data should be available in Mercury.
Radical: Writing from Mercury to disk array directly over fiber channel

## Slide 9: Things to do
By the end of October
Fastbus and VME over Ethernet
Fastbus and VME over cable.
Interface to online.
Updating Phat on CVS with final code.
More convenient handling of pedestals.
After
Overhaul of the L0/L1/L2 logic.
Implementation of the XLM module.

## Slide 10: Summary
Mercury can process 520 Hz safely. No FEC errors! Higher rate requires optimization of the system. Expect 700 Hz.
Sustained data taking rate (from Silicon to disk) for today = 28 MB, 280 Hz. Limited @ 28 MB/s by disk bottleneck @ 35 MB/s and Mercury-fpdp link @ 330 Hz. Can be improved to ~ 35 MB/s.
Compression does not work yet due to messed up pedestals. Need 1 day test with powered FECs.
Disk array optimization finished. Sustained writing speed limit = 42-50 MB/s. Simultaneous writing @ 42 MB/s and reading @ 28 MB/s.
Adding one more disk writing chain will increase Mercury-disk rate to 28*2 = 56 MB/s or 560 Hz. Guaranteed.
All DAQ code & designs are in CVS