## Slide 1: Event Manager Interconnections
CI 07
Done1 06
Veto1 05
Trig 1 04
CI 03
CI 02
ExtBusy 01
Trig0 00
PreSc0 12
Scaler0 11
Veto2 09
Trig 2 08
PreSc1 13
PreSc2 14
PreSc3 15
TrigVME ENG.CI
CO 07
CO 06
CO 05
CO 04
CO 03
CO 02
Done 01
Busy 00
TrigPulse 11
StrobCI 10
CO 09
CO 08
PreSc0 13
PreSc1 13
PreSc2 14
ENG.CO
Done2 10
PreSc3 15
CI 07
CI 06
CI 05
CI 04
CI 03
CI 02
CI 01
CI 00
CI 12
CI 11
CI 09
CI 08
MyTrig 13
CI 14
HandShk 15
CO 07
CO 06
CO 05
CO 04
CO 03
CO 02
CO 01
CO 00
CO 11
CO 10
CO 09
CO 08
CO 13
MyTrig 13
Veto 14
ING.CO
CI 10
Ack 15
Plastic ING.CI
CO 07
CO 06
CO 05
CO 04
CO 03
CO 02
CO 01
CO 00
CO 11
CO 10
CO 09
CO 08
CO 13
MyTrig 13
Veto 14
ING.CO
Ack 15
Silicon ING.CI
LED Trigger
Twisted pair color rule:
The ‘-’ wire is black!
- +
- +
- +
- +
- +
- +
CI 07
CI 06
CI 05
CI 04
CI 03
CI 02
CI 01
CI 00
CI 12
CI 11
CI 09
CI 08
MyTrig 13
CI 14
HandShk 15
CI 10

## Slide 2: TOF Gate Logic 12/99
794
GDG
4616
ECL-NIM-ECL
Trigger
TDC Gate to CAT.ICAS
Trigger to ENG.CI[0]
Delay cable 64 ns
Test pulses
for TDC
10 ns
Legend
794
GDG
Delay output
10 ns
ADC Gate, to CAT.ADCG
Veto from DAQ
Twisted pair wires,
The color represent the actual color of the +wire
NIM cables
Twisted pair color rule:
The ‘-’ wire is black!

## Slide 3: Run 2390. TDC resolution, TOF cable delays
TDC Gate to CAT.ICAS
462
Time
Calibrator
Start
Stop
NIM-ECL
365AL
Event
Manager
Trig
Busy
ECL-NIM
ECL-ECL
63 ns
ECL-ECL
6 ns
6 ns
Cable delay 78 ns
6 ns
ECL-ECL
63 ns
ECL-ECL
ECL-ECL
63 ns
ECL-ECL
NIM-ECL
63 ns
ECL-ECL
TDC 0948
TDC 0949
TDC 0950
TDC 0951
794
GDG
ECL-NIM
NIM-ECL
ADC Gate to CAT.ADCG
422
FO
Discriminators in the tunnel
All otherTDCs
8 ns
0 ns
52 ns
63 ns
Trig TDCs
6 ns
32 ns
diff=32 ns
The event manager is connected in normal mode: the main trigger is generated by event manager.

## Slide 4: Run 2390,2391. Results
In the run 2391 the TDC gate was taken directly from the AND gate 365AL by doing this we eliminate possible jitter in the event Manager
Comparing two tables we can conclude that :
The Event Manager does not affect the resolution of the TDC
The delay introduced by the Event Manager is 22 ns.
The best achievable TDC resolution is 100 ps in the range of 0-100ns, and 200-250 ps in the range of 100-800ns
The Round Trip Delay for TOF signals is 655-669 ns