## Slide 1: Proposal 1. Using time stretchers
1:20 Time Stretcher
1:20 Time Stretcher
2
40
800
T0A
T0B
SpecTrig Logic Module
Using two time stretchers STT100 stretch the coincidence signal at least 100 times. Measure width of the signal in STLM using internal 100 MHz clock. Using both edges of the clock it is possible to get SPLM time resolution of 5 ns.
Pros:
Should work
Occupies one input on the STLM
If stretcher works according to specs then it should be possible to achieve resolution better than 2 cm
Cons:
Need to build a simple (NIM?) board
STT100 Specifications:
Inputs and Outputs: 1 each, differential ECL.
External Clock Required for long base: 60 - 100 MHz.
Normal Stretch Ratio: 20:1.
Range: 0-32 us
LSB: 25 ps when used with a 500 ps LSB multi-hit TDC. Calibration method employs LeCroy propriety technique; patent applied for.
Operation Modes: Common Stop or Common Start.
Double pulse Resolution: < 800 ns with 61.5 MHz clock.
Packaging: Surface mount board, 1.5" x 3".
Power: +9 V, +5 V, -5 V, -2 V, -9 V; 2 W total

## Slide 2: Proposal 2. Shifted clock register
D
D
D
D
D
D
3 cm twisted pairs
T0A
T0B
Program part of the STLM as a n-bit register. Connect T0A to D inputs of the register, connect T0B to the clock inputs using delay line (twisted pair cable with 3 cm spaced connectors). Requires n bits to cover 2*n*3 cm vertex range. Output of the register will reflect time difference. For example 000111 means that T0B is three taps prior the front of T0A, 111000 means that T0B is three taps prior the end of T0A.
Pros:
Very easy to implement
Can be modeled now using PCDP board
Cons:
May not work. As pointed by Wojtek Skulski the main problem is flip-flop metastability (see http://www.onsemi.com/pub/Collateral/AN1504-D.PDF).
Too delicate to adjust
Very sensitive to the noise on the board
Occupies n+1 inputs of the STLM
Requires calibration

## Slide 3: Proposal 3. Cancelled
FADC LeCroy 4504 (proposed by Gerrit) probably is the best solution.