M. Friedl, H. Pernegger

Phobos Spectrometer 
Module Tests

Phobos Collaboration Meeting

BNL, Apr 16–18, 1999

Contents

- Introduction

- Calibration and Source Test

- Calibration Details

- Module Studies

- Signal

- Noise

- Signal-to-Noise

- Summary

Introduction

- Terminology:

1 or 2 hybrid(s) + 2 to 4 sensors = 1 module

“Left”
Hybrid

“Right”
Hybrid

Silicon Sensors

- Status:

Assembly of Phobos Spectrometer finished

(42 installed modules)

Calibration and Source Test

- 2 module tests:

Calibration (performed on all hybrids and on all modules)

DAQ

Voltage

Step

Generator

## Calibration
Capacitor

VA-HDR1

## Detector
Capacitance

- Source Test (performed on final modules)

VA-HDR1

DAQ

90Sr source

Si
Sensor

Trigger
Scintillator

Source particle

- Why calibration?

- Testing, immediately repairing bad bond wires etc. before any further assembly step

- Obtaining gain (calibration) of every single channel (spectrometer including spares: 68k)

- Find defect and noisy channels

- Why source test?

- Final module test, similar to beam operation

- Study module properties

- Determine signal and signal-to-noise

Measured data is filled into the Assembly Database (accessible via Phobos Web pages)

Calibration Details

- Step pulse over capacitor

- Q = C V

- Ex.: 1 mV on 2 pF gives 12500 e-

- Design considerations to minimize noise

- Output parameter: Channel gain [mV/fC]

- Calibration test setup

- Hybrid/Module

- Adapter board with step DAC and pulse generator board (E. Griesmayer, H. Frais-Kölbl), range 0..350 fC, diff. lin. better than 1%

- CERN Repeater cards (A. Rudge)

- NIM sequencer logic

- National Instruments ADC board

- PC running LabView

- Calibration components

- 1 capacitor is located close to each VA chip on the hybrids

- Switching logic is integrated in the VA chips

- Gain Uniformity

- Within 1 chip: <3% pp

- Chip-to-chip: <15% pp

- Calibration allows to indentify

- good channels

- pinholes

- channel shorts

- Calibration modes

- “Standard”: done on every hybrid and module, running for about 6 months

- “Precision”: done on selected modules, special effort on stability and parameter settings, done for module properties measurements (running for a few days)

- Calibration repeatability

- Standard: <3% (no effort to maintain voltages like FEC)

- Precision: <1% (as expected with FEC)

- No charge loss due to detector capacitance

- Could be calculated from calibrations of hybrid (w/o sensor) and module (w/ sensor) to see large effects >10%

- Detector capacitance: 25 pF (typ.)

- VA input capacitance: not measured

- Charge loss below standard mode error 
 VA input C >> detector C

- Sample gain map for MOD20001

- Comparison of average chip signals without and with channel calibration

Module Studies

- Levels of study

- Single sensor level

- Single module level

- Module/sensor comparison level

- Typical example: MOD20001

- 4 sensors (thicknesses 308..310 m)

- 5 rows, 400 columns

Origin

Corner

Columns

Rows

- Sensor level

- Sensor signal distribution (norm. to 300m)

- Fit function: Landau convoluted with Gaussian

3 parameters:

## Width, MP, Area

1 parameter:

Sigma

- Gaussian sigma represents electronic noise and intrinsic detector fluctuations, typically about 2x electronic noise

- Peak of convolute is slightly shifted from pure Landau peak (MP)

- Sensor RMS noise map

- Sensor MP signal vs. row

- Module level

- Signal and SNR distributions (norm. to 300m)

MP = 21523 e- = 77.9 keV

SNR = 16.4

- Relative MP signal vs. row

Uniformity better than 1% pp !!!

Signal

- Absolute MP signal vs. sensor type (normalized to 300 m thickness)

- Acquired with 90Sr  source

- Average source signal: 21081 e-

- Signal range: 20119..21552 e-

- Signal MP agrees with capacitive loss calculation (charge sharing between detector and VA input capacities)

- Relative MP signal vs. row

- Uniformity: <1%

- Very uniform even though readout lines vary in length (3..65 mm)

- Mystery

- 3% (1% at testbeam) difference between front and back half of hybrids

- Intrinsic module property, reason: ?

- We can exclude external effects because:

- Calibration and source tests are done with the same set of repeater boards

- Switching the ADC channels does not change anything

- Same (though smaller) effect in testbeam with totally different DAQ hard- and software

Noise

Single channel physical connections

RP

Sensor

+HV

Single channel preamp noise model

RP

RS

ILeak

TP

ENCC

C

| symbol | represents | typ. value |
| --- | --- | --- |
| ILeak | Fraction of sensor leakage current | 5 nA |
| RP | Parallel resistance (polysilicon) | 2 M |
| C | Capacitance of pad against backplane, between pads and between metal 1 and metal 2 layers | 25 pF |
| RS | Effective serial trace resistance | 40  |
| TP | Shaper peaking time | 1.2 s |
| ENCC | Preamp noise (k + C · d) | 950+5/pF e- |

- Total preamp equivalent noise charge

ENC2=ENCC2+ENCILeak2+ENCRP2+ENCRS2

- Contributions

ENCC=k+C d

ENCILeak=106ILeak TP

ENCRP=758TPRP

ENCRS=0.395 CRSTP

with ENC [e-], k [e-], d [e-/pF], 
ILeak [nA], TP [s], RP [M], RS []

- Preamp noise with typical values (as stated above)

| ENCC = | 1078 | e- |
| --- | --- | --- |
| ENCILeak = | 249 | e- |
| ENCRP = | 562 | e- |
| ENCRS = | 60 | e- |
| ENC = | 1243 | e- |

- Shaper/Buffer and readout noise adds to preamp noise

- Estimated shaper/buffer noise: 100 e-

- Measured test setup readout noise: 450 e-

- Calculated and measured noise vs. sensor type

- Calculation and measurement agree very well

- Measured RMS noise:

- Average noise: 1383 e-

- Noise range: 1230..1517 e-

Signal-to-Noise

- SNR vs. Sensor Type (normalized to 300 m)

- Measured SNR:

- Average SNR: 15.2

- SNR range: 14.4..16.9

Summary

- Stable calibration and source test setups

- Gain of every single channel written into assembly database

- Gain uniformity: 
<3% pp within chip, <15% pp chip-to-chip

- Calibration repeatability during tests
<1% precision

- Gain before and after sensor mounting is the same within error bars

- No significant charge loss by detector capacitance

- Due to large VA input capacitance

- Mystery

- 3% signal difference between hybrid halves

- “The general picture”

- Landau-Gaussian convolute fit function

- Typ. values (MOD20001):

- MP signal (90Sr source) = 21500 e-

- RMS noise = 1347 e-

- SNR = 16.0

- The sensor properties:

- Signal (normalized to 300 m thickness)

- Average MP signal: 21081 e-

- Signal range: 20119..21552 e-
(depending on sensor type)

- Uniformity: Signal vs. row <1% on all sensor types  very uniform!!!

- Noise

- Primarily depends on detector capacitance

- Average RMS noise: 1383 e-

- Noise range: 1230..1517 e-
(depending on sensor type)

- Agrees very well with model

- Signal-to-noise (normalized to 300 m thickness)

- Average SNR: 15.2

- SNR range: 14.4..16.9
(depending on sensor type)

P. Shulek et al., Sov. J. Nucl. Phys. 4, 400 (1967)